## Slide 1
OUT-OF-BOX EXTERNAL VALIDATION
Production Single-Lead AFib Detector
on PhysioNet/CinC Challenge 2017
Zero retraining · zero tuning · held-out test set
Technical & Clinical Executive Briefing
2026-06-02

## Slide 2
EXECUTIVE SUMMARY
Strong out-of-box AFib performance on an unseen device
88.2%
Sensitivity
AFib recordings caught
93.5%
Specificity
normal recordings left alone
0.95
ROC-AUC
overall discrimination
0.76
F1 score
balanced precision / recall
BOTTOM LINE
A detector calibrated only on our own ambulatory single-lead data generalizes to a different handheld device at 0.95 AUC. The production post-processing turns a raw model that fires indiscriminately (F1 0.50) into a deployable detector (F1 0.76) with clinically usable specificity.
ECGFounder · Out-of-Box AFib Validation on PhysioNet/CinC Challenge 2017
2

## Slide 3
MOTIVATION
Why validate on an external dataset?
Calibrated on one device
The production AFib head was tuned on the fzark ambulatory single-lead cohort. We must prove it holds on data it never saw.
An independent benchmark
Challenge 2017: 8,528 publicly-labeled single-lead recordings from a different handheld device (AliveCor, 300 Hz).
Zero-leakage generalization
The model was never trained on this data — a true out-of-box test of real-world robustness.
ECGFounder · Out-of-Box AFib Validation on PhysioNet/CinC Challenge 2017
3

## Slide 4
DATA & METHOD
Dataset, split, and recording-level detection
Challenge 2017 — 8,528 single-lead recordings (30–60 s)
Label | Meaning | n | Role
N | Normal sinus | 5,076 | neg (clean)
A | Atrial fibrillation | 758 | target
O | Other rhythm | 2,415 | context
~ | Noisy | 279 | context
Only A maps to a production scope event (AFib). N is the clean negative; O and ~ are reported for context only (mixed / noise), excluded from binary metrics.
STRATIFIED 80 / 10 / 10 SPLIT
Per-class, seed-fixed, frozen as a manifest. Validation = tuning context · Test = held-out report (852 recordings).
RECORDING-LEVEL DETECTION
Tile each recording into 10 s windows → preprocess (300→500 Hz, notch, band-pass, normalize) → score every window. A recording is flagged AFib if any window fires — identical to the production recording report.
ECGFounder · Out-of-Box AFib Validation on PhysioNet/CinC Challenge 2017
4

## Slide 5
THE PRODUCTION MODEL
A three-stage pipeline on a frozen foundation backbone
Backbone
ECGFounder 1-lead
Frozen 150-head foundation model. Raw per-event probabilities.
L1 · Calibration
fzarkSL projection
Re-calibrates data-rich heads (AFib, Brady, Tachy) to device operating points.
L2 · Arbiter
GT-matched rules
Deterministic suppression of impossible co-firings. Cuts false alarms, no learned weights.
This study isolates each stage's contribution to the final result.
ECGFounder · Out-of-Box AFib Validation on PhysioNet/CinC Challenge 2017
5

## Slide 6
PRIMARY RESULTS · HELD-OUT TEST
Raw head → L1 calibration → full production
Chart: 
| Raw head @0.5 | + L1 (L2-off) | Full production
Sensitivity | 0.987 | 0.895 | 0.882
Specificity | 0.708 | 0.876 | 0.935
F1 | 0.502 | 0.657 | 0.761
WHAT THE PIPELINE BUYS
Specificity
0.71 → 0.94
+23 pp false-positive reduction on normals
Precision (PPV)
0.34 → 0.67
nearly doubled
F1
0.50 → 0.76
undeployable → deployable
Sensitivity cost
0.99 → 0.88
modest, intentional trade
ECGFounder · Out-of-Box AFib Validation on PhysioNet/CinC Challenge 2017
6

## Slide 7
STAGE CONTRIBUTION
L2 arbitration: more specificity at almost no sensitivity cost
Chart: 
| Sensitivity (A) | Specificity (N) | Fire-rate Other ↓
L2-off (L1 only) | 0.895 | 0.876 | 0.282
L2-on (full production) | 0.882 | 0.935 | 0.178
Specificity
+5.9 pp
0.876 → 0.935
Other fire-rate
−10.4 pp
0.282 → 0.178
Sensitivity
−1.3 pp
0.895 → 0.882
ECGFounder · Out-of-Box AFib Validation on PhysioNet/CinC Challenge 2017
7

## Slide 8
KEY TECHNICAL FINDING
The overlay calibrates the decision point — not the ranking
Raw head ranks slightly better…
Raw head ROC-AUC 0.983 vs L1 0.946 (test).
…but the raw head is useless at its default 0.5 threshold — it fires on 29% of normal recordings.
Implication: high AUC ≠ deployable. A usable operating point is what matters in production.
WHAT THE OVERLAY DELIVERS
L1 maps the raw score to a usable threshold — specificity 0.876 at τ = 0.54.
L2 adds specificity on top → 0.935, changing the binary decision, not the score (so AUC is unchanged L2-off vs on).
Headroom: re-fitting L1 toward this distribution could recover ranking AUC for a multi-device operating point.
ECGFounder · Out-of-Box AFib Validation on PhysioNet/CinC Challenge 2017
8

## Slide 9
CLINICAL INTERPRETATION
What the numbers mean at the bedside
Catches ~8 of 9 AFib
88% sensitivity on an unseen handheld device — strong for ambulatory screening / triage.
Clears ~19 of 20 normals
93.5% specificity keeps the false-alarm burden low, reducing clinician alert fatigue.
Fewer spurious alerts
−23 pp false-positive reduction vs the raw model means far fewer non-AFib alerts reach review.
“Other” fires ~18%
Expected — that class contains AF-like and other arrhythmias; not a pure false-positive rate.
ECGFounder · Out-of-Box AFib Validation on PhysioNet/CinC Challenge 2017
9

## Slide 10
LIMITATIONS & NEXT STEPS
Honest scope and where we go from here
Limitations
Only AFib maps 1:1 to a production scope event — other arrhythmia heads not evaluable here.
“Other” / “Noisy” are not clean negatives — shown as context, excluded from binary metrics.
Recording-level any-window-fires rollup is sensitive to recording length.
Modest positive count (76 AFib in test) — consistency with validation mitigates this.
Next steps
Optionally re-fit L1 for a multi-device operating point to recover ranking-AUC headroom.
Extend validation to additional external single-lead datasets.
Report confidence intervals with a larger positive sample.
Stage remaining tp_rex classes to broaden multi-event coverage.
ECGFounder · Out-of-Box AFib Validation on PhysioNet/CinC Challenge 2017
10

## Slide 11
Conclusions
1
Generalizes out-of-box
0.95 AUC, 88% / 94% sensitivity / specificity on an independent single-lead device — the calibration is not overfit to our device.
2
The overlay is essential
Without L1 + L2 the raw model is undeployable (fires on 29% of normals). L2 alone buys +5.9 pp specificity at ~zero recall cost.
3
Clear path forward
Re-fit L1 for multi-device deployment; broaden external validation; report CIs on a larger positive sample.
Artifacts: PROD_AFIB.md · CHALLENGE2017_STUDY_REPORT.md · csv/challenge2017_split.csv · scripts/eval_challenge2017_prod.py